Storing Memories

Up to nine front panel control settings can be stored for recall at a

later time as follows:

1.  To store a set-up:

After adjusting the instrument to the desired settings, press the

STORE push button and the desired storage location number from 1 to 9

using the numeric keypad.  The number selected will flash in the

MEASUREMENT/ENTRY display.  Note that if a set-up was previously

stored at the location selected (1-9), the old setting will be erased

and the new set up will be stored at that location.

2.  To recall a set-up:

Press the RECALL push button and then the desired setup number from 1

to 9 using the numeric keypad.  The set-up stored at that location

will be displayed and become immediately active.

3.  To reset the instrument to the turn-on condition:

Press RECALL and then 0.